THE TAPR TICC:
Can This One Weird Trick Reduce the Noise Floor?

The TDC7200 chip used in the TICC timestamping counter builds its measurement results from three components, made using two very different counters. One counter counts edges of the 100NS clock derived from the external 10 MHz reference clock. The other counts the number of loops through a 63 stage, free-running, ring oscillator with a nominal period of 55 ps.

As shown in the diagram (copied from the TDC7200 data sheet), the measurement process is:

1. On START, count number of ring oscillator periods to next 100NS edge (time1).
2. From that edge, count 100NS ticks until the first edge after STOP (clock_count1).
3. On STOP, count to next 100NS edge (time2).
4. Time of Flight (TOF) is time1 + clock_count1 – time2.

Now, here's the thing. In the TICC, the STOP pulse is actually the 100 kHz COARSE_CLOCK edge passed through a gate system. COARSE_CLOCK is derived from 100NS through a PIC divider. The timing relationship between the two is based on the propagation delay through the divider and the hardware logic on the board. So, the time from the COARSE_CLOCK risign edge applied to the TDC STOP pin until the next 100NS edge is fixed. With current TICC hardware and software, oscilloscope measurements show that time to be about 68 ns.

So: if the relationship between the STOP pulse and the 100NS clock edge is fixed, is it necessary to measure that time in each measurement? Or could a fixed value be used instead? Would doing so reduce system noise

By putting the TICC into "Debug" mode, you can see the raw data registers output by the TDC7200 chip. The "time2Result" field contains the number of ring oscillator cycles between STOP and the next 100NS edge.

Data from a "loopback" test using coherent 10 MHz and 1 PPS sources is shown below. Deduce from it what you may!

Debug Results – Raw Register Values (n=1000)
	time1	time2	clock1	cal1	cal2	calculated tof
mean	848.210	1270.898	1000	1839.23	36822.822	9.997699777e-05
min	845	1270	1000	1836	36767	9.9976796e-05
max	851	1272	1000	1842	36866	9.9977141e-05
range	6	2	0	6	99	3.45e-10
sdev	1.0113	0.4688	0	0.7910	11.9706	5.987e-11

Values generated by TVB's "stat.c" program from ticc_rev_d_loopback_chA_debug.txt (available from https://github.com/TAPR/TICC/docs) generated by a TICC Rev. D board running software version 20170215.1. Data captured on 20 Feb. 2016 and in the accompanying file ticc_rev_d_loopback_chA_debug.txt. 10 MHz reference was IsoTemp 131-100 OXCO, chA input was 1 PPS from on-board PICdiv. Field order is time1, time2, clock1, cal1, cal2, PICstop, tof, timestamp.

000848 001271 001000 001839 036830 73240178 0.000099976974 7324.017700023026 chA
000848 001271 001000 001839 036826 73250178 0.000099976972 7325.017700023028 chA
000848 001271 001000 001839 036820 73260178 0.000099976968 7326.017700023032 chA
000849 001271 001000 001840 036821 73270178 0.000099977022 7327.017700022978 chA
000850 001271 001000 001840 036829 73280178 0.000099977082 7328.017700022918 chA
000849 001271 001000 001839 036825 73290178 0.000099977025 7329.017700022975 chA
000848 001270 001000 001839 036814 73300178 0.000099977018 7330.017700022982 chA
000849 001271 001000 001839 036810 73310178 0.000099977015 7331.017700022985 chA
000849 001270 001000 001839 036821 73320178 0.000099977077 7332.017700022923 chA
000848 001271 001000 001838 036813 73330178 0.000099976964 7333.017700023036 chA
000849 001271 001000 001839 036818 73340178 0.000099977021 7334.017700022979 chA
000849 001271 001000 001839 036826 73350178 0.000099977026 7335.017700022974 chA
000849 001271 001000 001839 036813 73360178 0.000099977017 7336.017700022983 chA
000849 001271 001000 001840 036846 73370178 0.000099977038 7337.017700022962 chA
000849 001270 001000 001839 036811 73380178 0.000099977071 7338.017700022929 chA
000848 001271 001000 001839 036814 73390178 0.000099976964 7339.017700023036 chA
000849 001270 001000 001839 036813 73400178 0.000099977072 7340.017700022928 chA
000849 001271 001000 001840 036839 73410178 0.000099977034 7341.017700022966 chA
000848 001271 001000 001841 036846 73420178 0.000099976983 7342.017700023017 chA
000848 001271 001000 001840 036827 73430178 0.000099976972 7343.017700023028 chA
000849 001271 001000 001839 036820 73440178 0.000099977022 7344.017700022978 chA
000850 001271 001000 001839 036818 73450178 0.000099977075 7345.017700022925 chA
000849 001271 001000 001839 036829 73460178 0.000099977028 7346.017700022972 chA
000849 001270 001000 001840 036836 73470178 0.000099977086 7347.017700022914 chA
000849 001271 001000 001838 036815 73480178 0.000099977019 7348.017700022981 chA
000849 001271 001000 001839 036825 73490178 0.000099977025 7349.017700022975 chA
000849 001271 001000 001840 036838 73500178 0.000099977033 7350.017700022967 chA
000850 001271 001000 001839 036819 73510178 0.000099977076 7351.017700022924 chA
000849 001271 001000 001840 036843 73520178 0.000099977036 7352.017700022964 chA
000848 001271 001000 001838 036802 73530178 0.000099976956 7353.017700023044 chA
000849 001271 001000 001839 036828 73540178 0.000099977027 7354.017700022973 chA
000850 001271 001000 001839 036815 73550178 0.000099977073 7355.017700022927 chA
000849 001271 001000 001841 036841 73560178 0.000099977035 7356.017700022965 chA
000849 001271 001000 001840 036824 73570178 0.000099977024 7357.017700022976 chA
000848 001271 001000 001839 036822 73580178 0.000099976969 7358.017700023031 chA
000849 001271 001000 001839 036817 73590178 0.000099977020 7359.017700022980 chA
000850 001271 001000 001840 036829 73600178 0.000099977082 7360.017700022918 chA
000849 001271 001000 001841 036838 73610178 0.000099977033 7361.017700022967 chA
000848 001271 001000 001839 036818 73620178 0.000099976966 7362.017700023034 chA
000849 001271 001000 001839 036825 73630178 0.000099977025 7363.017700022975 chA
000848 001271 001000 001840 036826 73640178 0.000099976971 7364.017700023029 chA
000850 001271 001000 001840 036838 73650178 0.000099977088 7365.017700022912 chA
000848 001271 001000 001839 036819 73660178 0.000099976967 7366.017700023033 chA
000848 001271 001000 001838 036807 73670178 0.000099976960 7367.017700023040 chA
000849 001271 001000 001840 036829 73680178 0.000099977027 7368.017700022973 chA
000849 001271 001000 001839 036821 73690178 0.000099977023 7369.017700022977 chA
000849 001271 001000 001840 036831 73700178 0.000099977029 7370.017700022971 chA
000849 001271 001000 001840 036827 73710178 0.000099977026 7371.017700022974 chA
000849 001271 001000 001840 036831 73720178 0.000099977029 7372.017700022971 chA
000849 001271 001000 001839 036822 73730178 0.000099977023 7373.017700022977 chA
000849 001271 001000 001839 036820 73740178 0.000099977022 7374.017700022978 chA
000849 001270 001000 001839 036808 73750178 0.000099977069 7375.017700022931 chA
000848 001271 001000 001839 036824 73760178 0.000099976970 7376.017700023030 chA
000849 001270 001000 001839 036820 73770178 0.000099977076 7377.017700022924 chA
000850 001271 001000 001840 036823 73780178 0.000099977078 7378.017700022922 chA
000849 001271 001000 001839 036810 73790178 0.000099977015 7379.017700022985 chA
000848 001270 001000 001839 036812 73800178 0.000099977017 7380.017700022983 chA
000849 001271 001000 001839 036827 73810178 0.000099977027 7381.017700022973 chA
000849 001271 001000 001840 036838 73820178 0.000099977033 7382.017700022967 chA
000849 001271 001000 001838 036809 73830178 0.000099977015 7383.017700022985 chA
000849 001270 001000 001838 036803 73840178 0.000099977066 7384.017700022934 chA
000849 001271 001000 001838 036808 73850178 0.000099977015 7385.017700022985 chA
000848 001272 001000 001838 036816 73860178 0.000099976911 7386.017700023089 chA
000849 001271 001000 001838 036801 73870178 0.000099977010 7387.017700022990 chA
000848 001271 001000 001839 036826 73880178 0.000099976972 7388.017700023028 chA
000848 001271 001000 001840 036835 73890178 0.000099976977 7389.017700023023 chA
000850 001271 001000 001840 036827 73900178 0.000099977080 7390.017700022920 chA
000850 001271 001000 001839 036806 73910178 0.000099977067 7391.017700022933 chA
000850 001271 001000 001839 036827 73920178 0.000099977081 7392.017700022919 chA
000849 001271 001000 001839 036823 73930178 0.000099977024 7393.017700022976 chA
000848 001271 001000 001839 036824 73940178 0.000099976970 7394.017700023030 chA
000849 001271 001000 001840 036834 73950178 0.000099977031 7395.017700022969 chA
000850 001271 001000 001839 036819 73960178 0.000099977076 7396.017700022924 chA
000849 001270 001000 001838 036799 73970178 0.000099977063 7397.017700022937 chA
000848 001271 001000 001839 036820 73980178 0.000099976968 7398.017700023032 chA
000850 001271 001000 001840 036834 73990178 0.000099977085 7399.017700022915 chA
000850 001271 001000 001839 036820 74000178 0.000099977076 7400.017700022924 chA
000849 001271 001000 001839 036820 74010178 0.000099977022 7401.017700022978 chA
000849 001271 001000 001840 036833 74020178 0.000099977030 7402.017700022970 chA
000849 001272 001000 001838 036817 74030178 0.000099976966 7403.017700023034 chA
000849 001271 001000 001839 036823 74040178 0.000099977024 7404.017700022976 chA
000849 001271 001000 001839 036818 74050178 0.000099977021 7405.017700022979 chA
000849 001271 001000 001840 036825 74060178 0.000099977025 7406.017700022975 chA
000848 001271 001000 001839 036829 74070178 0.000099976974 7407.017700023026 chA
000848 001271 001000 001841 036849 74080178 0.000099976985 7408.017700023015 chA
000850 001270 001000 001839 036821 74090178 0.000099977132 7409.017700022868 chA
000849 001271 001000 001840 036834 74100178 0.000099977031 7410.017700022969 chA
000848 001271 001000 001840 036835 74110178 0.000099976977 7411.017700023023 chA
000849 001270 001000 001839 036825 74120178 0.000099977080 7412.017700022920 chA
000850 001271 001000 001841 036850 74130178 0.000099977095 7413.017700022905 chA
000849 001271 001000 001838 036798 74140178 0.000099977008 7414.017700022992 chA
000848 001271 001000 001839 036817 74150178 0.000099976966 7415.017700023034 chA
000848 001271 001000 001840 036833 74160178 0.000099976975 7416.017700023025 chA
000849 001271 001000 001839 036823 74170178 0.000099977024 7417.017700022976 chA
000849 001271 001000 001838 036810 74180178 0.000099977016 7418.017700022984 chA
000849 001271 001000 001840 036843 74190178 0.000099977036 7419.017700022964 chA
000849 001270 001000 001839 036822 74200178 0.000099977078 7420.017700022922 chA
000848 001271 001000 001840 036825 74210178 0.000099976970 7421.017700023030 chA
000849 001271 001000 001840 036836 74220178 0.000099977032 7422.017700022968 chA
000850 001271 001000 001839 036820 74230178 0.000099977076 7423.017700022924 chA
000849 001271 001000 001840 036828 74240178 0.000099977027 7424.017700022973 chA
000849 001270 001000 001840 036836 74250178 0.000099977086 7425.017700022914 chA
000849 001271 001000 001840 036836 74260178 0.000099977032 7426.017700022968 chA
000850 001271 001000 001840 036837 74270178 0.000099977087 7427.017700022913 chA
000850 001271 001000 001840 036824 74280178 0.000099977078 7428.017700022922 chA
000848 001271 001000 001840 036824 74290178 0.000099976970 7429.017700023030 chA
000850 001270 001000 001839 036822 74300178 0.000099977132 7430.017700022868 chA
000850 001271 001000 001841 036852 74310178 0.000099977096 7431.017700022904 chA
000850 001271 001000 001839 036823 74320178 0.000099977078 7432.017700022922 chA
000849 001271 001000 001838 036813 74330178 0.000099977018 7433.017700022982 chA
000848 001271 001000 001839 036823 74340178 0.000099976970 7434.017700023030 chA
000849 001271 001000 001840 036843 74350178 0.000099977036 7435.017700022964 chA
000849 001271 001000 001839 036820 74360178 0.000099977022 7436.017700022978 chA
000849 001271 001000 001839 036822 74370178 0.000099977023 7437.017700022977 chA
000848 001271 001000 001839 036825 74380178 0.000099976971 7438.017700023029 chA
000849 001271 001000 001840 036828 74390178 0.000099977027 7439.017700022973 chA
000850 001271 001000 001839 036829 74400178 0.000099977082 7440.017700022918 chA
000849 001271 001000 001838 036808 74410178 0.000099977015 7441.017700022985 chA
000848 001271 001000 001838 036810 74420178 0.000099976962 7442.017700023038 chA
000848 001271 001000 001839 036823 74430178 0.000099976970 7443.017700023030 chA
000848 001271 001000 001839 036820 74440178 0.000099976968 7444.017700023032 chA
000850 001272 001000 001839 036827 74450178 0.000099977027 7445.017700022973 chA
000850 001271 001000 001838 036801 74460178 0.000099977065 7446.017700022935 chA
000848 001271 001000 001839 036822 74470178 0.000099976969 7447.017700023031 chA
000850 001271 001000 001839 036815 74480178 0.000099977073 7448.017700022927 chA
000849 001271 001000 001839 036815 74490178 0.000099977019 7449.017700022981 chA
000849 001272 001000 001839 036827 74500178 0.000099976972 7450.017700023028 chA
000848 001271 001000 001840 036841 74510178 0.000099976981 7451.017700023019 chA
000849 001271 001000 001839 036822 74520178 0.000099977023 7452.017700022977 chA
000849 001271 001000 001838 036802 74530178 0.000099977011 7453.017700022989 chA
000850 001271 001000 001839 036824 74540178 0.000099977079 7454.017700022921 chA
000848 001271 001000 001839 036825 74550178 0.000099976971 7455.017700023029 chA
000848 001271 001000 001839 036822 74560178 0.000099976969 7456.017700023031 chA
000847 001271 001000 001841 036845 74570178 0.000099976928 7457.017700023072 chA
000849 001271 001000 001839 036823 74580178 0.000099977024 7458.017700022976 chA
000849 001271 001000 001839 036807 74590178 0.000099977013 7459.017700022987 chA
000849 001270 001000 001840 036826 74600178 0.000099977080 7460.017700022920 chA
000850 001271 001000 001839 036813 74610178 0.000099977072 7461.017700022928 chA
000850 001271 001000 001838 036792 74620178 0.000099977059 7462.017700022941 chA
000850 001272 001000 001840 036835 74630178 0.000099977031 7463.017700022969 chA
000848 001270 001000 001837 036790 74640178 0.000099977004 7464.017700022996 chA
000848 001271 001000 001840 036829 74650178 0.000099976973 7465.017700023027 chA
000849 001272 001000 001841 036841 74660178 0.000099976980 7466.017700023020 chA
000850 001271 001000 001839 036812 74670178 0.000099977071 7467.017700022929 chA
000849 001272 001000 001838 036816 74680178 0.000099976966 7468.017700023034 chA
000849 001272 001000 001839 036831 74690178 0.000099976975 7469.017700023025 chA
000849 001270 001000 001839 036828 74700178 0.000099977082 7470.017700022918 chA
000849 001272 001000 001840 036830 74710178 0.000099976974 7471.017700023026 chA
000850 001272 001000 001839 036823 74720178 0.000099977024 7472.017700022976 chA
000849 001272 001000 001838 036818 74730178 0.000099976967 7473.017700023033 chA
000849 001271 001000 001840 036839 74740178 0.000099977034 7474.017700022966 chA
000849 001271 001000 001840 036835 74750178 0.000099977031 7475.017700022969 chA
000850 001271 001000 001839 036826 74760178 0.000099977080 7476.017700022920 chA
000849 001271 001000 001840 036835 74770178 0.000099977031 7477.017700022969 chA
000847 001271 001000 001839 036812 74780178 0.000099976908 7478.017700023092 chA
000848 001271 001000 001840 036834 74790178 0.000099976976 7479.017700023024 chA
000849 001271 001000 001840 036836 74800178 0.000099977032 7480.017700022968 chA
000849 001271 001000 001840 036823 74810178 0.000099977023 7481.017700022977 chA
000849 001271 001000 001838 036799 74820178 0.000099977009 7482.017700022991 chA
000847 001270 001000 001839 036805 74830178 0.000099976958 7483.017700023042 chA
000849 001270 001000 001839 036813 74840178 0.000099977072 7484.017700022928 chA
000849 001271 001000 001839 036822 74850178 0.000099977023 7485.017700022977 chA
000849 001272 001000 001839 036825 74860178 0.000099976971 7486.017700023029 chA
000849 001271 001000 001839 036819 74870178 0.000099977021 7487.017700022979 chA
000849 001272 001000 001839 036826 74880178 0.000099976972 7488.017700023028 chA
000849 001271 001000 001838 036805 74890178 0.000099977013 7489.017700022987 chA
000850 001271 001000 001839 036821 74900178 0.000099977077 7490.017700022923 chA
000849 001271 001000 001840 036832 74910178 0.000099977029 7491.017700022971 chA
000848 001270 001000 001839 036818 74920178 0.000099977021 7492.017700022979 chA
000848 001271 001000 001839 036823 74930178 0.000099976970 7493.017700023030 chA
000849 001272 001000 001837 036770 74940178 0.000099976936 7494.017700023064 chA
000849 001271 001000 001840 036827 74950178 0.000099977026 7495.017700022974 chA
000850 001271 001000 001839 036828 74960178 0.000099977082 7496.017700022918 chA
000849 001271 001000 001839 036816 74970178 0.000099977019 7497.017700022981 chA
000849 001271 001000 001839 036817 74980178 0.000099977020 7498.017700022980 chA
000849 001271 001000 001839 036823 74990178 0.000099977024 7499.017700022976 chA
000848 001271 001000 001839 036826 75000178 0.000099976972 7500.017700023028 chA
000848 001270 001000 001840 036832 75010178 0.000099977029 7501.017700022971 chA
000849 001271 001000 001838 036802 75020178 0.000099977011 7502.017700022989 chA
000850 001271 001000 001839 036823 75030178 0.000099977078 7503.017700022922 chA
000850 001271 001000 001838 036814 75040178 0.000099977073 7504.017700022927 chA
000848 001271 001000 001840 036829 75050178 0.000099976973 7505.017700023027 chA
000848 001271 001000 001840 036826 75060178 0.000099976971 7506.017700023029 chA
000848 001271 001000 001840 036835 75070178 0.000099976977 7507.017700023023 chA
000850 001271 001000 001840 036827 75080178 0.000099977080 7508.017700022920 chA
000847 001271 001000 001838 036804 75090178 0.000099976903 7509.017700023097 chA
000847 001271 001000 001840 036839 75100178 0.000099976925 7510.017700023075 chA
000848 001272 001000 001838 036817 75110178 0.000099976912 7511.017700023088 chA
000849 001271 001000 001839 036825 75120178 0.000099977025 7512.017700022975 chA
000849 001270 001000 001840 036842 75130178 0.000099977090 7513.017700022910 chA
000850 001270 001000 001840 036832 75140178 0.000099977138 7514.017700022862 chA
000849 001270 001000 001839 036817 75150178 0.000099977075 7515.017700022925 chA
000849 001271 001000 001840 036831 75160178 0.000099977029 7516.017700022971 chA
000849 001271 001000 001840 036844 75170178 0.000099977037 7517.017700022963 chA
000849 001271 001000 001839 036822 75180178 0.000099977023 7518.017700022977 chA
000849 001271 001000 001841 036846 75190178 0.000099977038 7519.017700022962 chA
000848 001271 001000 001839 036832 75200178 0.000099976975 7520.017700023025 chA
000850 001272 001000 001841 036848 75210178 0.000099977039 7521.017700022961 chA
000849 001271 001000 001839 036824 75220178 0.000099977025 7522.017700022975 chA
000849 001271 001000 001840 036823 75230178 0.000099977023 7523.017700022977 chA
000847 001272 001000 001838 036809 75240178 0.000099976852 7524.017700023148 chA
000850 001271 001000 001840 036832 75250178 0.000099977084 7525.017700022916 chA
000850 001271 001000 001839 036827 75260178 0.000099977081 7526.017700022919 chA
000849 001271 001000 001840 036831 75270178 0.000099977029 7527.017700022971 chA
000848 001271 001000 001840 036828 75280178 0.000099976972 7528.017700023028 chA
000849 001272 001000 001838 036819 75290178 0.000099976968 7529.017700023032 chA
000849 001271 001000 001840 036838 75300178 0.000099977033 7530.017700022967 chA
000849 001271 001000 001839 036828 75310178 0.000099977027 7531.017700022973 chA
000849 001271 001000 001838 036807 75320178 0.000099977014 7532.017700022986 chA
000848 001272 001000 001839 036830 75330178 0.000099976920 7533.017700023080 chA
000849 001271 001000 001839 036820 75340178 0.000099977022 7534.017700022978 chA
000849 001272 001000 001838 036812 75350178 0.000099976963 7535.017700023037 chA
000850 001271 001000 001839 036823 75360178 0.000099977078 7536.017700022922 chA
000848 001271 001000 001840 036832 75370178 0.000099976975 7537.017700023025 chA
000849 001271 001000 001839 036820 75380178 0.000099977022 7538.017700022978 chA
000848 001270 001000 001839 036805 75390178 0.000099977012 7539.017700022988 chA
000851 001271 001000 001840 036837 75400178 0.000099977141 7540.017700022859 chA
000849 001271 001000 001838 036803 75410178 0.000099977012 7541.017700022988 chA
000848 001271 001000 001839 036811 75420178 0.000099976962 7542.017700023038 chA
000849 001271 001000 001840 036839 75430178 0.000099977034 7543.017700022966 chA
000850 001271 001000 001840 036843 75440178 0.000099977091 7544.017700022909 chA
000849 001271 001000 001840 036834 75450178 0.000099977031 7545.017700022969 chA
000849 001271 001000 001840 036823 75460178 0.000099977023 7546.017700022977 chA
000848 001271 001000 001840 036833 75470178 0.000099976975 7547.017700023025 chA
000849 001271 001000 001841 036854 75480178 0.000099977043 7548.017700022957 chA
000849 001271 001000 001839 036810 75490178 0.000099977015 7549.017700022985 chA
000849 001271 001000 001839 036826 75500178 0.000099977026 7550.017700022974 chA
000849 001270 001000 001838 036811 75510178 0.000099977071 7551.017700022929 chA
000848 001271 001000 001839 036815 75520178 0.000099976964 7552.017700023036 chA
000850 001271 001000 001839 036826 75530178 0.000099977080 7553.017700022920 chA
000849 001271 001000 001839 036823 75540178 0.000099977024 7554.017700022976 chA
000849 001271 001000 001839 036820 75550178 0.000099977022 7555.017700022978 chA
000849 001270 001000 001840 036837 75560178 0.000099977087 7556.017700022913 chA
000850 001271 001000 001839 036817 75570178 0.000099977075 7557.017700022925 chA
000849 001272 001000 001840 036835 75580178 0.000099976977 7558.017700023023 chA
000849 001272 001000 001839 036829 75590178 0.000099976974 7559.017700023026 chA
000848 001271 001000 001839 036835 75600178 0.000099976977 7560.017700023023 chA
000849 001272 001000 001841 036860 75610178 0.000099976993 7561.017700023007 chA
000850 001271 001000 001840 036829 75620178 0.000099977082 7562.017700022918 chA
000849 001272 001000 001840 036835 75630178 0.000099976977 7563.017700023023 chA
000848 001271 001000 001840 036831 75640178 0.000099976974 7564.017700023026 chA
000849 001271 001000 001840 036830 75650178 0.000099977028 7565.017700022972 chA
000849 001270 001000 001840 036823 75660178 0.000099977078 7566.017700022922 chA
000849 001271 001000 001840 036832 75670178 0.000099977029 7567.017700022971 chA
000848 001271 001000 001838 036803 75680178 0.000099976957 7568.017700023043 chA
000849 001271 001000 001839 036821 75690178 0.000099977023 7569.017700022977 chA
000848 001271 001000 001839 036822 75700178 0.000099976969 7570.017700023031 chA
000850 001271 001000 001839 036825 75710178 0.000099977080 7571.017700022920 chA
000848 001271 001000 001839 036826 75720178 0.000099976972 7572.017700023028 chA
000848 001271 001000 001838 036805 75730178 0.000099976958 7573.017700023042 chA
000849 001270 001000 001839 036822 75740178 0.000099977078 7574.017700022922 chA
000850 001271 001000 001839 036827 75750178 0.000099977081 7575.017700022919 chA
000849 001271 001000 001840 036824 75760178 0.000099977024 7576.017700022976 chA
000849 001271 001000 001840 036826 75770178 0.000099977025 7577.017700022975 chA
000848 001272 001000 001840 036837 75780178 0.000099976924 7578.017700023076 chA
000850 001271 001000 001840 036841 75790178 0.000099977090 7579.017700022910 chA
000849 001271 001000 001838 036797 75800178 0.000099977008 7580.017700022992 chA
000849 001271 001000 001838 036807 75810178 0.000099977014 7581.017700022986 chA
000848 001271 001000 001841 036845 75820178 0.000099976983 7582.017700023017 chA
000850 001271 001000 001839 036825 75830178 0.000099977080 7583.017700022920 chA
000849 001271 001000 001840 036842 75840178 0.000099977036 7584.017700022964 chA
000850 001271 001000 001839 036823 75850178 0.000099977078 7585.017700022922 chA
000849 001271 001000 001839 036812 75860178 0.000099977017 7586.017700022983 chA
000849 001272 001000 001840 036832 75870178 0.000099976975 7587.017700023025 chA
000849 001271 001000 001840 036838 75880178 0.000099977033 7588.017700022967 chA
000848 001271 001000 001839 036817 75890178 0.000099976966 7589.017700023034 chA
000849 001271 001000 001839 036825 75900178 0.000099977025 7590.017700022975 chA
000848 001271 001000 001840 036830 75910178 0.000099976974 7591.017700023026 chA
000849 001271 001000 001839 036823 75920178 0.000099977024 7592.017700022976 chA
000849 001272 001000 001841 036854 75930178 0.000099976989 7593.017700023011 chA
000849 001271 001000 001840 036838 75940178 0.000099977033 7594.017700022967 chA
000847 001270 001000 001840 036837 75950178 0.000099976978 7595.017700023022 chA
000849 001271 001000 001839 036823 75960178 0.000099977024 7596.017700022976 chA
000850 001270 001000 001840 036832 75970178 0.000099977138 7597.017700022862 chA
000850 001272 001000 001839 036820 75980178 0.000099977022 7598.017700022978 chA
000849 001271 001000 001840 036831 75990178 0.000099977029 7599.017700022971 chA
000848 001272 001000 001839 036823 76000178 0.000099976915 7600.017700023085 chA
000849 001271 001000 001840 036836 76010178 0.000099977032 7601.017700022968 chA
000849 001271 001000 001839 036822 76020178 0.000099977023 7602.017700022977 chA
000849 001272 001000 001839 036836 76030178 0.000099976978 7603.017700023022 chA
000849 001271 001000 001840 036823 76040178 0.000099977023 7604.017700022977 chA
000849 001271 001000 001840 036829 76050178 0.000099977027 7605.017700022973 chA
000849 001271 001000 001840 036834 76060178 0.000099977031 7606.017700022969 chA
000849 001271 001000 001839 036818 76070178 0.000099977021 7607.017700022979 chA
000849 001271 001000 001839 036820 76080178 0.000099977022 7608.017700022978 chA
000849 001271 001000 001839 036820 76090178 0.000099977022 7609.017700022978 chA
000849 001271 001000 001839 036815 76100178 0.000099977019 7610.017700022981 chA
000849 001271 001000 001839 036827 76110178 0.000099977027 7611.017700022973 chA
000849 001272 001000 001840 036843 76120178 0.000099976982 7612.017700023018 chA
000847 001271 001000 001839 036812 76130178 0.000099976908 7613.017700023092 chA
000849 001271 001000 001839 036823 76140178 0.000099977024 7614.017700022976 chA
000850 001271 001000 001839 036823 76150178 0.000099977078 7615.017700022922 chA
000850 001271 001000 001840 036831 76160178 0.000099977083 7616.017700022917 chA
000849 001270 001000 001840 036824 76170178 0.000099977078 7617.017700022922 chA
000848 001271 001000 001838 036814 76180178 0.000099976964 7618.017700023036 chA
000849 001271 001000 001840 036835 76190178 0.000099977031 7619.017700022969 chA
000849 001270 001000 001839 036810 76200178 0.000099977070 7620.017700022930 chA
000850 001271 001000 001840 036831 76210178 0.000099977083 7621.017700022917 chA
000848 001271 001000 001840 036830 76220178 0.000099976974 7622.017700023026 chA
000847 001271 001000 001838 036812 76230178 0.000099976909 7623.017700023091 chA
000850 001271 001000 001839 036824 76240178 0.000099977079 7624.017700022921 chA
000850 001270 001000 001839 036816 76250178 0.000099977128 7625.017700022872 chA
000849 001271 001000 001839 036833 76260178 0.000099977031 7626.017700022969 chA
000848 001271 001000 001841 036846 76270178 0.000099976983 7627.017700023017 chA
000849 001271 001000 001840 036838 76280178 0.000099977033 7628.017700022967 chA
000849 001271 001000 001839 036818 76290178 0.000099977021 7629.017700022979 chA
000850 001271 001000 001839 036817 76300178 0.000099977075 7630.017700022925 chA
000848 001271 001000 001839 036819 76310178 0.000099976967 7631.017700023033 chA
000849 001271 001000 001840 036842 76320178 0.000099977036 7632.017700022964 chA
000849 001271 001000 001839 036822 76330178 0.000099977023 7633.017700022977 chA
000850 001271 001000 001838 036808 76340178 0.000099977069 7634.017700022931 chA
000849 001271 001000 001838 036805 76350178 0.000099977013 7635.017700022987 chA
000847 001271 001000 001838 036813 76360178 0.000099976909 7636.017700023091 chA
000849 001271 001000 001839 036817 76370178 0.000099977020 7637.017700022980 chA
000850 001271 001000 001839 036833 76380178 0.000099977085 7638.017700022915 chA
000848 001270 001000 001839 036823 76390178 0.000099977024 7639.017700022976 chA
000847 001271 001000 001840 036834 76400178 0.000099976922 7640.017700023078 chA
000850 001271 001000 001840 036849 76410178 0.000099977095 7641.017700022905 chA
000850 001271 001000 001840 036825 76420178 0.000099977079 7642.017700022921 chA
000849 001271 001000 001839 036817 76430178 0.000099977020 7643.017700022980 chA
000848 001271 001000 001838 036802 76440178 0.000099976956 7644.017700023044 chA
000849 001271 001000 001839 036817 76450178 0.000099977020 7645.017700022980 chA
000848 001271 001000 001840 036839 76460178 0.000099976979 7646.017700023021 chA
000850 001271 001000 001840 036838 76470178 0.000099977088 7647.017700022912 chA
000849 001272 001000 001839 036823 76480178 0.000099976970 7648.017700023030 chA
000848 001271 001000 001840 036832 76490178 0.000099976975 7649.017700023025 chA
000847 001271 001000 001839 036823 76500178 0.000099976915 7650.017700023085 chA
000849 001271 001000 001840 036837 76510178 0.000099977033 7651.017700022967 chA
000850 001272 001000 001838 036814 76520178 0.000099977019 7652.017700022981 chA
000849 001271 001000 001840 036824 76530178 0.000099977024 7653.017700022976 chA
000848 001271 001000 001839 036821 76540178 0.000099976968 7654.017700023032 chA
000848 001271 001000 001839 036819 76550178 0.000099976967 7655.017700023033 chA
000850 001271 001000 001840 036832 76560178 0.000099977084 7656.017700022916 chA
000849 001271 001000 001839 036828 76570178 0.000099977027 7657.017700022973 chA
000849 001271 001000 001839 036818 76580178 0.000099977021 7658.017700022979 chA
000847 001271 001000 001839 036809 76590178 0.000099976906 7659.017700023094 chA
000849 001271 001000 001839 036817 76600178 0.000099977020 7660.017700022980 chA
000850 001271 001000 001839 036827 76610178 0.000099977081 7661.017700022919 chA
000849 001271 001000 001839 036827 76620178 0.000099977027 7662.017700022973 chA
000849 001271 001000 001839 036818 76630178 0.000099977021 7663.017700022979 chA
000850 001271 001000 001839 036822 76640178 0.000099977078 7664.017700022922 chA
000850 001271 001000 001839 036817 76650178 0.000099977075 7665.017700022925 chA
000849 001271 001000 001840 036828 76660178 0.000099977027 7666.017700022973 chA
000848 001271 001000 001839 036818 76670178 0.000099976966 7667.017700023034 chA
000847 001270 001000 001839 036820 76680178 0.000099976968 7668.017700023032 chA
000848 001271 001000 001839 036834 76690178 0.000099976977 7669.017700023023 chA
000849 001270 001000 001840 036830 76700178 0.000099977082 7670.017700022918 chA
000848 001271 001000 001840 036833 76710178 0.000099976975 7671.017700023025 chA
000848 001271 001000 001840 036831 76720178 0.000099976974 7672.017700023026 chA
000848 001270 001000 001839 036820 76730178 0.000099977022 7673.017700022978 chA
000849 001271 001000 001840 036838 76740178 0.000099977033 7674.017700022967 chA
000848 001271 001000 001839 036816 76750178 0.000099976965 7675.017700023035 chA
000849 001271 001000 001840 036829 76760178 0.000099977027 7676.017700022973 chA
000848 001271 001000 001840 036837 76770178 0.000099976978 7677.017700023022 chA
000849 001270 001000 001840 036822 76780178 0.000099977077 7678.017700022923 chA
000850 001272 001000 001839 036824 76790178 0.000099977025 7679.017700022975 chA
000849 001271 001000 001839 036827 76800178 0.000099977027 7680.017700022973 chA
000847 001271 001000 001839 036813 76810178 0.000099976909 7681.017700023091 chA
000848 001271 001000 001840 036840 76820178 0.000099976980 7682.017700023020 chA
000848 001271 001000 001839 036822 76830178 0.000099976969 7683.017700023031 chA
000849 001271 001000 001840 036826 76840178 0.000099977025 7684.017700022975 chA
000848 001272 001000 001840 036837 76850178 0.000099976924 7685.017700023076 chA
000848 001271 001000 001839 036823 76860178 0.000099976970 7686.017700023030 chA
000849 001271 001000 001840 036844 76870178 0.000099977037 7687.017700022963 chA
000849 001270 001000 001840 036833 76880178 0.000099977084 7688.017700022916 chA
000849 001271 001000 001840 036832 76890178 0.000099977029 7689.017700022971 chA
000849 001271 001000 001840 036834 76900178 0.000099977031 7690.017700022969 chA
000849 001270 001000 001840 036836 76910178 0.000099977086 7691.017700022914 chA
000850 001271 001000 001840 036838 76920178 0.000099977088 7692.017700022912 chA
000849 001271 001000 001840 036828 76930178 0.000099977027 7693.017700022973 chA
000847 001272 001000 001839 036829 76940178 0.000099976865 7694.017700023135 chA
000847 001271 001000 001840 036832 76950178 0.000099976920 7695.017700023080 chA
000849 001271 001000 001840 036841 76960178 0.000099977035 7696.017700022965 chA
000849 001271 001000 001839 036817 76970178 0.000099977020 7697.017700022980 chA
000849 001271 001000 001839 036812 76980178 0.000099977017 7698.017700022983 chA
000847 001271 001000 001839 036828 76990178 0.000099976918 7699.017700023082 chA
000849 001270 001000 001838 036803 77000178 0.000099977066 7700.017700022934 chA
000849 001271 001000 001836 036774 77010178 0.000099976994 7701.017700023006 chA
000849 001271 001000 001840 036826 77020178 0.000099977025 7702.017700022975 chA
000848 001271 001000 001838 036803 77030178 0.000099976957 7703.017700023043 chA
000848 001271 001000 001839 036823 77040178 0.000099976970 7704.017700023030 chA
000849 001271 001000 001838 036795 77050178 0.000099977006 7705.017700022994 chA
000850 001272 001000 001838 036804 77060178 0.000099977012 7706.017700022988 chA
000850 001271 001000 001840 036836 77070178 0.000099977086 7707.017700022914 chA
000849 001271 001000 001838 036810 77080178 0.000099977016 7708.017700022984 chA
000849 001271 001000 001840 036835 77090178 0.000099977031 7709.017700022969 chA
000849 001272 001000 001839 036822 77100178 0.000099976969 7710.017700023031 chA
000850 001271 001000 001841 036842 77110178 0.000099977090 7711.017700022910 chA
000849 001271 001000 001838 036800 77120178 0.000099977010 7712.017700022990 chA
000849 001270 001000 001839 036821 77130178 0.000099977077 7713.017700022923 chA
000849 001271 001000 001840 036824 77140178 0.000099977024 7714.017700022976 chA
000850 001271 001000 001839 036816 77150178 0.000099977074 7715.017700022926 chA
000849 001271 001000 001839 036818 77160178 0.000099977021 7716.017700022979 chA
000847 001271 001000 001840 036834 77170178 0.000099976922 7717.017700023078 chA
000849 001271 001000 001839 036825 77180178 0.000099977025 7718.017700022975 chA
000849 001270 001000 001839 036823 77190178 0.000099977078 7719.017700022922 chA
000850 001271 001000 001838 036804 77200178 0.000099977067 7720.017700022933 chA
000848 001271 001000 001839 036811 77210178 0.000099976962 7721.017700023038 chA
000847 001271 001000 001841 036835 77220178 0.000099976922 7722.017700023078 chA
000849 001271 001000 001839 036830 77230178 0.000099977029 7723.017700022971 chA
000849 001272 001000 001840 036829 77240178 0.000099976973 7724.017700023027 chA
000849 001272 001000 001840 036836 77250178 0.000099976977 7725.017700023023 chA
000848 001270 001000 001839 036819 77260178 0.000099977021 7726.017700022979 chA
000849 001271 001000 001840 036833 77270178 0.000099977030 7727.017700022970 chA
000850 001271 001000 001839 036822 77280178 0.000099977078 7728.017700022922 chA
000850 001272 001000 001840 036838 77290178 0.000099977033 7729.017700022967 chA
000849 001271 001000 001839 036823 77300178 0.000099977024 7730.017700022976 chA
000848 001271 001000 001839 036817 77310178 0.000099976966 7731.017700023034 chA
000848 001271 001000 001840 036846 77320178 0.000099976984 7732.017700023016 chA
000849 001271 001000 001840 036826 77330178 0.000099977025 7733.017700022975 chA
000850 001271 001000 001839 036823 77340178 0.000099977078 7734.017700022922 chA
000848 001271 001000 001839 036823 77350178 0.000099976970 7735.017700023030 chA
000848 001271 001000 001840 036842 77360178 0.000099976981 7736.017700023019 chA
000849 001271 001000 001840 036828 77370178 0.000099977027 7737.017700022973 chA
000850 001271 001000 001839 036822 77380178 0.000099977078 7738.017700022922 chA
000849 001271 001000 001840 036830 77390178 0.000099977028 7739.017700022972 chA
000848 001271 001000 001842 036866 77400178 0.000099976996 7740.017700023004 chA
000849 001271 001000 001840 036829 77410178 0.000099977027 7741.017700022973 chA
000850 001271 001000 001839 036827 77420178 0.000099977081 7742.017700022919 chA
000849 001271 001000 001839 036819 77430178 0.000099977021 7743.017700022979 chA
000847 001271 001000 001840 036839 77440178 0.000099976925 7744.017700023075 chA
000848 001271 001000 001839 036816 77450178 0.000099976965 7745.017700023035 chA
000850 001271 001000 001840 036837 77460178 0.000099977087 7746.017700022913 chA
000850 001271 001000 001838 036813 77470178 0.000099977073 7747.017700022927 chA
000847 001271 001000 001838 036807 77480178 0.000099976905 7748.017700023095 chA
000848 001271 001000 001838 036813 77490178 0.000099976964 7749.017700023036 chA
000849 001271 001000 001839 036832 77500178 0.000099977030 7750.017700022970 chA
000849 001271 001000 001839 036824 77510178 0.000099977025 7751.017700022975 chA
000847 001270 001000 001839 036808 77520178 0.000099976960 7752.017700023040 chA
000849 001270 001000 001839 036824 77530178 0.000099977079 7753.017700022921 chA
000848 001271 001000 001839 036830 77540178 0.000099976974 7754.017700023026 chA
000849 001271 001000 001838 036797 77550178 0.000099977008 7755.017700022992 chA
000848 001271 001000 001839 036815 77560178 0.000099976964 7756.017700023036 chA
000848 001271 001000 001838 036804 77570178 0.000099976958 7757.017700023042 chA
000847 001271 001000 001840 036838 77580178 0.000099976924 7758.017700023076 chA
000849 001270 001000 001840 036839 77590178 0.000099977088 7759.017700022912 chA
000849 001270 001000 001840 036834 77600178 0.000099977085 7760.017700022915 chA
000849 001271 001000 001839 036833 77610178 0.000099977031 7761.017700022969 chA
000848 001271 001000 001839 036818 77620178 0.000099976966 7762.017700023034 chA
000848 001271 001000 001839 036819 77630178 0.000099976967 7763.017700023033 chA
000848 001271 001000 001839 036826 77640178 0.000099976972 7764.017700023028 chA
000849 001271 001000 001839 036826 77650178 0.000099977026 7765.017700022974 chA
000849 001271 001000 001839 036830 77660178 0.000099977029 7766.017700022971 chA
000848 001271 001000 001839 036822 77670178 0.000099976969 7767.017700023031 chA
000849 001271 001000 001839 036827 77680178 0.000099977027 7768.017700022973 chA
000850 001271 001000 001840 036829 77690178 0.000099977082 7769.017700022918 chA
000847 001271 001000 001840 036829 77700178 0.000099976918 7770.017700023082 chA
000847 001271 001000 001839 036828 77710178 0.000099976918 7771.017700023082 chA
000847 001271 001000 001839 036819 77720178 0.000099976912 7772.017700023088 chA
000849 001270 001000 001839 036817 77730178 0.000099977075 7773.017700022925 chA
000849 001271 001000 001840 036829 77740178 0.000099977027 7774.017700022973 chA
000848 001271 001000 001840 036831 77750178 0.000099976974 7775.017700023026 chA
000848 001271 001000 001838 036796 77760178 0.000099976952 7776.017700023048 chA
000849 001271 001000 001838 036814 77770178 0.000099977019 7777.017700022981 chA
000850 001271 001000 001839 036815 77780178 0.000099977073 7778.017700022927 chA
000849 001270 001000 001838 036795 77790178 0.000099977061 7779.017700022939 chA
000848 001270 001000 001839 036825 77800178 0.000099977025 7780.017700022975 chA
000848 001270 001000 001839 036818 77810178 0.000099977021 7781.017700022979 chA
000849 001271 001000 001839 036823 77820178 0.000099977024 7782.017700022976 chA
000848 001271 001000 001839 036823 77830178 0.000099976970 7783.017700023030 chA
000848 001271 001000 001840 036839 77840178 0.000099976979 7784.017700023021 chA
000847 001271 001000 001838 036809 77850178 0.000099976907 7785.017700023093 chA
000849 001271 001000 001841 036838 77860178 0.000099977033 7786.017700022967 chA
000850 001270 001000 001840 036823 77870178 0.000099977132 7787.017700022868 chA
000849 001270 001000 001838 036800 77880178 0.000099977064 7788.017700022936 chA
000848 001270 001000 001839 036825 77890178 0.000099977025 7789.017700022975 chA
000847 001270 001000 001840 036831 77900178 0.000099976974 7790.017700023026 chA
000849 001272 001000 001839 036816 77910178 0.000099976965 7791.017700023035 chA
000849 001271 001000 001840 036829 77920178 0.000099977027 7792.017700022973 chA
000848 001271 001000 001839 036824 77930178 0.000099976970 7793.017700023030 chA
000848 001271 001000 001840 036835 77940178 0.000099976977 7794.017700023023 chA
000849 001271 001000 001840 036830 77950178 0.000099977028 7795.017700022972 chA
000849 001270 001000 001840 036827 77960178 0.000099977080 7796.017700022920 chA
000849 001271 001000 001839 036829 77970178 0.000099977028 7797.017700022972 chA
000847 001271 001000 001840 036841 77980178 0.000099976926 7798.017700023074 chA
000847 001271 001000 001838 036807 77990178 0.000099976905 7799.017700023095 chA
000848 001271 001000 001838 036812 78000178 0.000099976963 7800.017700023037 chA
000848 001271 001000 001839 036825 78010178 0.000099976971 7801.017700023029 chA
000848 001271 001000 001839 036823 78020178 0.000099976970 7802.017700023030 chA
000848 001271 001000 001840 036835 78030178 0.000099976977 7803.017700023023 chA
000849 001271 001000 001840 036835 78040178 0.000099977031 7804.017700022969 chA
000849 001271 001000 001840 036838 78050178 0.000099977033 7805.017700022967 chA
000847 001271 001000 001839 036818 78060178 0.000099976912 7806.017700023088 chA
000847 001271 001000 001840 036834 78070178 0.000099976922 7807.017700023078 chA
000847 001271 001000 001839 036830 78080178 0.000099976920 7808.017700023080 chA
000849 001271 001000 001839 036815 78090178 0.000099977019 7809.017700022981 chA
000850 001270 001000 001839 036827 78100178 0.000099977136 7810.017700022864 chA
000849 001271 001000 001840 036841 78110178 0.000099977035 7811.017700022965 chA
000847 001270 001000 001840 036829 78120178 0.000099976973 7812.017700023027 chA
000849 001271 001000 001840 036829 78130178 0.000099977027 7813.017700022973 chA
000848 001271 001000 001840 036832 78140178 0.000099976975 7814.017700023025 chA
000847 001271 001000 001839 036826 78150178 0.000099976917 7815.017700023083 chA
000846 001271 001000 001839 036820 78160178 0.000099976859 7816.017700023141 chA
000846 001271 001000 001839 036816 78170178 0.000099976856 7817.017700023144 chA
000848 001270 001000 001839 036810 78180178 0.000099977015 7818.017700022985 chA
000849 001270 001000 001839 036810 78190178 0.000099977070 7819.017700022930 chA
000848 001271 001000 001840 036828 78200178 0.000099976972 7820.017700023028 chA
000847 001270 001000 001840 036832 78210178 0.000099976975 7821.017700023025 chA
000849 001271 001000 001839 036823 78220178 0.000099977024 7822.017700022976 chA
000850 001271 001000 001839 036824 78230178 0.000099977079 7823.017700022921 chA
000848 001271 001000 001840 036839 78240178 0.000099976979 7824.017700023021 chA
000846 001271 001000 001839 036819 78250178 0.000099976858 7825.017700023142 chA
000847 001271 001000 001839 036818 78260178 0.000099976912 7826.017700023088 chA
000848 001271 001000 001840 036829 78270178 0.000099976973 7827.017700023027 chA
000848 001270 001000 001840 036826 78280178 0.000099977025 7828.017700022975 chA
000847 001271 001000 001840 036834 78290178 0.000099976922 7829.017700023078 chA
000848 001270 001000 001840 036824 78300178 0.000099977024 7830.017700022976 chA
000848 001272 001000 001839 036831 78310178 0.000099976920 7831.017700023080 chA
000849 001270 001000 001841 036848 78320178 0.000099977094 7832.017700022906 chA
000847 001271 001000 001839 036826 78330178 0.000099976917 7833.017700023083 chA
000847 001271 001000 001839 036825 78340178 0.000099976916 7834.017700023084 chA
000848 001271 001000 001838 036805 78350178 0.000099976958 7835.017700023042 chA
000849 001271 001000 001838 036805 78360178 0.000099977013 7836.017700022987 chA
000848 001271 001000 001838 036795 78370178 0.000099976952 7837.017700023048 chA
000846 001270 001000 001839 036822 78380178 0.000099976914 7838.017700023086 chA
000847 001271 001000 001838 036804 78390178 0.000099976903 7839.017700023097 chA
000847 001270 001000 001840 036826 78400178 0.000099976971 7840.017700023029 chA
000847 001271 001000 001839 036823 78410178 0.000099976915 7841.017700023085 chA
000847 001271 001000 001839 036825 78420178 0.000099976916 7842.017700023084 chA
000847 001271 001000 001840 036845 78430178 0.000099976929 7843.017700023071 chA
000847 001271 001000 001839 036825 78440178 0.000099976916 7844.017700023084 chA
000847 001271 001000 001839 036818 78450178 0.000099976912 7845.017700023088 chA
000847 001271 001000 001839 036820 78460178 0.000099976913 7846.017700023087 chA
000847 001271 001000 001841 036849 78470178 0.000099976931 7847.017700023069 chA
000847 001271 001000 001840 036832 78480178 0.000099976920 7848.017700023080 chA
000847 001271 001000 001839 036817 78490178 0.000099976911 7849.017700023089 chA
000847 001270 001000 001839 036819 78500178 0.000099976967 7850.017700023033 chA
000847 001271 001000 001839 036814 78510178 0.000099976909 7851.017700023091 chA
000848 001271 001000 001839 036819 78520178 0.000099976967 7852.017700023033 chA
000847 001271 001000 001838 036803 78530178 0.000099976903 7853.017700023097 chA
000849 001271 001000 001839 036823 78540178 0.000099977024 7854.017700022976 chA
000847 001271 001000 001838 036805 78550178 0.000099976904 7855.017700023096 chA
000845 001271 001000 001839 036825 78560178 0.000099976808 7856.017700023192 chA
000846 001270 001000 001839 036827 78570178 0.000099976918 7857.017700023082 chA
000848 001271 001000 001839 036810 78580178 0.000099976961 7858.017700023039 chA
000849 001271 001000 001839 036812 78590178 0.000099977017 7859.017700022983 chA
000848 001271 001000 001838 036820 78600178 0.000099976968 7860.017700023032 chA
000846 001271 001000 001839 036815 78610178 0.000099976855 7861.017700023145 chA
000846 001271 001000 001838 036801 78620178 0.000099976847 7862.017700023153 chA
000848 001271 001000 001840 036833 78630178 0.000099976975 7863.017700023025 chA
000848 001271 001000 001840 036829 78640178 0.000099976973 7864.017700023027 chA
000847 001270 001000 001840 036831 78650178 0.000099976974 7865.017700023026 chA
000846 001271 001000 001839 036822 78660178 0.000099976860 7866.017700023140 chA
000847 001271 001000 001838 036804 78670178 0.000099976903 7867.017700023097 chA
000849 001271 001000 001839 036824 78680178 0.000099977025 7868.017700022975 chA
000849 001271 001000 001840 036833 78690178 0.000099977030 7869.017700022970 chA
000848 001271 001000 001840 036828 78700178 0.000099976972 7870.017700023028 chA
000847 001271 001000 001840 036837 78710178 0.000099976924 7871.017700023076 chA
000847 001272 001000 001839 036825 78720178 0.000099976862 7872.017700023138 chA
000849 001271 001000 001839 036823 78730178 0.000099977024 7873.017700022976 chA
000849 001271 001000 001840 036828 78740178 0.000099977027 7874.017700022973 chA
000847 001271 001000 001839 036818 78750178 0.000099976912 7875.017700023088 chA
000849 001271 001000 001840 036834 78760178 0.000099977031 7876.017700022969 chA
000848 001271 001000 001839 036811 78770178 0.000099976962 7877.017700023038 chA
000848 001271 001000 001840 036839 78780178 0.000099976979 7878.017700023021 chA
000848 001271 001000 001839 036829 78790178 0.000099976974 7879.017700023026 chA
000848 001270 001000 001838 036802 78800178 0.000099977011 7880.017700022989 chA
000849 001271 001000 001839 036819 78810178 0.000099977021 7881.017700022979 chA
000848 001271 001000 001840 036829 78820178 0.000099976973 7882.017700023027 chA
000847 001271 001000 001839 036814 78830178 0.000099976909 7883.017700023091 chA
000846 001270 001000 001840 036825 78840178 0.000099976916 7884.017700023084 chA
000847 001271 001000 001839 036822 78850178 0.000099976914 7885.017700023086 chA
000848 001270 001000 001839 036822 78860178 0.000099977023 7886.017700022977 chA
000849 001271 001000 001839 036822 78870178 0.000099977023 7887.017700022977 chA
000847 001271 001000 001839 036814 78880178 0.000099976909 7888.017700023091 chA
000847 001271 001000 001838 036802 78890178 0.000099976902 7889.017700023098 chA
000847 001270 001000 001840 036828 78900178 0.000099976972 7890.017700023028 chA
000848 001271 001000 001840 036827 78910178 0.000099976972 7891.017700023028 chA
000848 001271 001000 001839 036820 78920178 0.000099976968 7892.017700023032 chA
000847 001271 001000 001840 036829 78930178 0.000099976918 7893.017700023082 chA
000848 001271 001000 001838 036810 78940178 0.000099976962 7894.017700023038 chA
000849 001270 001000 001840 036822 78950178 0.000099977077 7895.017700022923 chA
000848 001271 001000 001838 036808 78960178 0.000099976960 7896.017700023040 chA
000847 001271 001000 001839 036826 78970178 0.000099976917 7897.017700023083 chA
000845 001271 001000 001839 036830 78980178 0.000099976811 7898.017700023189 chA
000847 001270 001000 001839 036826 78990178 0.000099976972 7899.017700023028 chA
000848 001271 001000 001839 036814 79000178 0.000099976964 7900.017700023036 chA
000848 001270 001000 001840 036836 79010178 0.000099977032 7901.017700022968 chA
000847 001270 001000 001840 036830 79020178 0.000099976974 7902.017700023026 chA
000849 001271 001000 001839 036824 79030178 0.000099977025 7903.017700022975 chA
000849 001271 001000 001839 036825 79040178 0.000099977025 7904.017700022975 chA
000850 001270 001000 001839 036812 79050178 0.000099977126 7905.017700022874 chA
000849 001271 001000 001840 036833 79060178 0.000099977030 7906.017700022970 chA
000847 001271 001000 001839 036817 79070178 0.000099976911 7907.017700023089 chA
000848 001271 001000 001840 036823 79080178 0.000099976969 7908.017700023031 chA
000849 001270 001000 001839 036812 79090178 0.000099977071 7909.017700022929 chA
000848 001271 001000 001841 036837 79100178 0.000099976977 7910.017700023023 chA
000845 001271 001000 001841 036857 79110178 0.000099976827 7911.017700023173 chA
000846 001271 001000 001839 036818 79120178 0.000099976857 7912.017700023143 chA
000848 001271 001000 001839 036821 79130178 0.000099976968 7913.017700023032 chA
000848 001270 001000 001840 036833 79140178 0.000099977030 7914.017700022970 chA
000848 001270 001000 001838 036809 79150178 0.000099977015 7915.017700022985 chA
000846 001271 001000 001839 036818 79160178 0.000099976857 7916.017700023143 chA
000848 001271 001000 001839 036823 79170178 0.000099976970 7917.017700023030 chA
000850 001271 001000 001840 036839 79180178 0.000099977088 7918.017700022912 chA
000847 001271 001000 001839 036816 79190178 0.000099976910 7919.017700023090 chA
000847 001271 001000 001838 036804 79200178 0.000099976903 7920.017700023097 chA
000848 001271 001000 001839 036813 79210178 0.000099976963 7921.017700023037 chA
000847 001271 001000 001838 036808 79220178 0.000099976906 7922.017700023094 chA
000849 001270 001000 001840 036828 79230178 0.000099977081 7923.017700022919 chA
000849 001271 001000 001838 036810 79240178 0.000099977016 7924.017700022984 chA
000847 001271 001000 001840 036834 79250178 0.000099976922 7925.017700023078 chA
000847 001271 001000 001840 036835 79260178 0.000099976922 7926.017700023078 chA
000849 001270 001000 001841 036854 79270178 0.000099977097 7927.017700022903 chA
000849 001270 001000 001839 036815 79280178 0.000099977073 7928.017700022927 chA
000847 001271 001000 001837 036791 79290178 0.000099976895 7929.017700023105 chA
000847 001270 001000 001839 036814 79300178 0.000099976964 7930.017700023036 chA
000849 001271 001000 001839 036824 79310178 0.000099977025 7931.017700022975 chA
000848 001270 001000 001840 036814 79320178 0.000099977017 7932.017700022983 chA
000846 001271 001000 001839 036812 79330178 0.000099976853 7933.017700023147 chA
000847 001271 001000 001839 036823 79340178 0.000099976915 7934.017700023085 chA
000847 001271 001000 001840 036831 79350178 0.000099976920 7935.017700023080 chA
000847 001271 001000 001839 036827 79360178 0.000099976918 7936.017700023082 chA
000848 001271 001000 001840 036833 79370178 0.000099976975 7937.017700023025 chA
000847 001271 001000 001840 036826 79380178 0.000099976916 7938.017700023084 chA
000847 001271 001000 001839 036820 79390178 0.000099976913 7939.017700023087 chA
000847 001270 001000 001839 036827 79400178 0.000099976972 7940.017700023028 chA
000850 001270 001000 001840 036833 79410178 0.000099977139 7941.017700022861 chA
000847 001271 001000 001839 036820 79420178 0.000099976913 7942.017700023087 chA
000847 001270 001000 001839 036810 79430178 0.000099976961 7943.017700023039 chA
000847 001270 001000 001839 036817 79440178 0.000099976966 7944.017700023034 chA
000849 001270 001000 001839 036805 79450178 0.000099977067 7945.017700022933 chA
000849 001270 001000 001839 036811 79460178 0.000099977071 7946.017700022929 chA
000847 001271 001000 001839 036827 79470178 0.000099976918 7947.017700023082 chA
000846 001271 001000 001840 036829 79480178 0.000099976864 7948.017700023136 chA
000848 001271 001000 001839 036819 79490178 0.000099976967 7949.017700023033 chA
000850 001271 001000 001840 036838 79500178 0.000099977088 7950.017700022912 chA
000848 001271 001000 001839 036817 79510178 0.000099976966 7951.017700023034 chA
000846 001270 001000 001838 036796 79520178 0.000099976898 7952.017700023102 chA
000847 001271 001000 001839 036815 79530178 0.000099976910 7953.017700023090 chA
000849 001271 001000 001841 036846 79540178 0.000099977038 7954.017700022962 chA
000849 001271 001000 001838 036810 79550178 0.000099977016 7955.017700022984 chA
000849 001271 001000 001839 036816 79560178 0.000099977019 7956.017700022981 chA
000847 001271 001000 001840 036830 79570178 0.000099976919 7957.017700023081 chA
000847 001271 001000 001838 036805 79580178 0.000099976904 7958.017700023096 chA
000847 001272 001000 001839 036827 79590178 0.000099976863 7959.017700023137 chA
000847 001271 001000 001839 036817 79600178 0.000099976911 7960.017700023089 chA
000848 001271 001000 001840 036835 79610178 0.000099976977 7961.017700023023 chA
000847 001271 001000 001840 036834 79620178 0.000099976922 7962.017700023078 chA
000848 001270 001000 001839 036817 79630178 0.000099977020 7963.017700022980 chA
000848 001270 001000 001839 036815 79640178 0.000099977019 7964.017700022981 chA
000848 001270 001000 001839 036819 79650178 0.000099977021 7965.017700022979 chA
000848 001271 001000 001839 036822 79660178 0.000099976969 7966.017700023031 chA
000849 001270 001000 001840 036827 79670178 0.000099977080 7967.017700022920 chA
000849 001271 001000 001839 036815 79680178 0.000099977019 7968.017700022981 chA
000849 001271 001000 001839 036827 79690178 0.000099977027 7969.017700022973 chA
000847 001271 001000 001840 036830 79700178 0.000099976919 7970.017700023081 chA
000847 001271 001000 001839 036820 79710178 0.000099976913 7971.017700023087 chA
000848 001272 001000 001838 036817 79720178 0.000099976912 7972.017700023088 chA
000848 001271 001000 001840 036838 79730178 0.000099976979 7973.017700023021 chA
000847 001270 001000 001839 036808 79740178 0.000099976960 7974.017700023040 chA
000848 001271 001000 001840 036835 79750178 0.000099976977 7975.017700023023 chA
000849 001270 001000 001839 036816 79760178 0.000099977074 7976.017700022926 chA
000848 001271 001000 001839 036814 79770178 0.000099976964 7977.017700023036 chA
000847 001270 001000 001840 036825 79780178 0.000099976970 7978.017700023030 chA
000846 001272 001000 001838 036807 79790178 0.000099976796 7979.017700023204 chA
000847 001271 001000 001839 036828 79800178 0.000099976918 7980.017700023082 chA
000849 001271 001000 001839 036817 79810178 0.000099977020 7981.017700022980 chA
000849 001270 001000 001840 036828 79820178 0.000099977081 7982.017700022919 chA
000847 001271 001000 001839 036818 79830178 0.000099976912 7983.017700023088 chA
000848 001272 001000 001840 036828 79840178 0.000099976918 7984.017700023082 chA
000848 001271 001000 001838 036809 79850178 0.000099976961 7985.017700023039 chA
000848 001271 001000 001840 036831 79860178 0.000099976974 7986.017700023026 chA
000849 001271 001000 001839 036821 79870178 0.000099977023 7987.017700022977 chA
000847 001271 001000 001839 036819 79880178 0.000099976912 7988.017700023088 chA
000848 001272 001000 001838 036809 79890178 0.000099976907 7989.017700023093 chA
000847 001271 001000 001839 036825 79900178 0.000099976916 7990.017700023084 chA
000849 001270 001000 001840 036827 79910178 0.000099977080 7991.017700022920 chA
000850 001272 001000 001839 036819 79920178 0.000099977021 7992.017700022979 chA
000846 001270 001000 001840 036823 79930178 0.000099976914 7993.017700023086 chA
000849 001271 001000 001840 036838 79940178 0.000099977033 7994.017700022967 chA
000848 001271 001000 001840 036826 79950178 0.000099976971 7995.017700023029 chA
000850 001270 001000 001839 036805 79960178 0.000099977121 7996.017700022879 chA
000848 001271 001000 001839 036812 79970178 0.000099976962 7997.017700023038 chA
000847 001271 001000 001839 036811 79980178 0.000099976907 7998.017700023093 chA
000848 001271 001000 001840 036834 79990178 0.000099976976 7999.017700023024 chA
000848 001271 001000 001840 036833 80000178 0.000099976975 8000.017700023025 chA
000849 001271 001000 001838 036808 80010178 0.000099977015 8001.017700022985 chA
000848 001271 001000 001840 036831 80020178 0.000099976974 8002.017700023026 chA
000847 001271 001000 001839 036824 80030178 0.000099976916 8003.017700023084 chA
000849 001271 001000 001838 036814 80040178 0.000099977019 8004.017700022981 chA
000849 001271 001000 001840 036846 80050178 0.000099977038 8005.017700022962 chA
000848 001271 001000 001839 036816 80060178 0.000099976965 8006.017700023035 chA
000847 001271 001000 001838 036800 80070178 0.000099976901 8007.017700023099 chA
000847 001272 001000 001838 036817 80080178 0.000099976857 8008.017700023143 chA
000848 001271 001000 001840 036825 80090178 0.000099976970 8009.017700023030 chA
000849 001271 001000 001839 036824 80100178 0.000099977025 8010.017700022975 chA
000849 001271 001000 001840 036836 80110178 0.000099977032 8011.017700022968 chA
000847 001271 001000 001839 036820 80120178 0.000099976913 8012.017700023087 chA
000848 001271 001000 001839 036823 80130178 0.000099976970 8013.017700023030 chA
000849 001270 001000 001837 036774 80140178 0.000099977048 8014.017700022952 chA
000848 001271 001000 001839 036816 80150178 0.000099976965 8015.017700023035 chA
000848 001270 001000 001838 036810 80160178 0.000099977016 8016.017700022984 chA
000848 001271 001000 001839 036825 80170178 0.000099976971 8017.017700023029 chA
000848 001271 001000 001838 036815 80180178 0.000099976965 8018.017700023035 chA
000850 001271 001000 001839 036823 80190178 0.000099977078 8019.017700022922 chA
000849 001272 001000 001839 036827 80200178 0.000099976972 8020.017700023028 chA
000847 001272 001000 001838 036800 80210178 0.000099976846 8021.017700023154 chA
000846 001271 001000 001840 036824 80220178 0.000099976861 8022.017700023139 chA
000850 001270 001000 001840 036825 80230178 0.000099977134 8023.017700022866 chA
000849 001270 001000 001839 036815 80240178 0.000099977073 8024.017700022927 chA
000847 001271 001000 001839 036825 80250178 0.000099976916 8025.017700023084 chA
000848 001270 001000 001838 036811 80260178 0.000099977017 8026.017700022983 chA
000849 001271 001000 001839 036822 80270178 0.000099977023 8027.017700022977 chA
000849 001271 001000 001839 036822 80280178 0.000099977023 8028.017700022977 chA
000848 001271 001000 001840 036836 80290178 0.000099976977 8029.017700023023 chA
000847 001271 001000 001839 036823 80300178 0.000099976915 8030.017700023085 chA
000849 001271 001000 001840 036832 80310178 0.000099977029 8031.017700022971 chA
000850 001271 001000 001839 036827 80320178 0.000099977081 8032.017700022919 chA
000849 001271 001000 001839 036820 80330178 0.000099977022 8033.017700022978 chA
000848 001271 001000 001839 036823 80340178 0.000099976970 8034.017700023030 chA
000847 001270 001000 001840 036826 80350178 0.000099976971 8035.017700023029 chA
000849 001271 001000 001838 036797 80360178 0.000099977008 8036.017700022992 chA
000850 001271 001000 001839 036821 80370178 0.000099977077 8037.017700022923 chA
000849 001271 001000 001841 036849 80380178 0.000099977040 8038.017700022960 chA
000848 001271 001000 001839 036819 80390178 0.000099976967 8039.017700023033 chA
000848 001270 001000 001840 036829 80400178 0.000099977027 8040.017700022973 chA
000849 001271 001000 001839 036828 80410178 0.000099977027 8041.017700022973 chA
000850 001271 001000 001839 036822 80420178 0.000099977078 8042.017700022922 chA
000848 001271 001000 001839 036817 80430178 0.000099976966 8043.017700023034 chA
000848 001271 001000 001839 036829 80440178 0.000099976974 8044.017700023026 chA
000847 001271 001000 001839 036813 80450178 0.000099976909 8045.017700023091 chA
000849 001270 001000 001840 036820 80460178 0.000099977076 8046.017700022924 chA
000850 001271 001000 001840 036832 80470178 0.000099977084 8047.017700022916 chA
000849 001271 001000 001840 036834 80480178 0.000099977031 8048.017700022969 chA
000847 001270 001000 001839 036823 80490178 0.000099976970 8049.017700023030 chA
000847 001271 001000 001839 036826 80500178 0.000099976917 8050.017700023083 chA
000849 001271 001000 001840 036828 80510178 0.000099977027 8051.017700022973 chA
000849 001271 001000 001839 036810 80520178 0.000099977015 8052.017700022985 chA
000848 001271 001000 001838 036803 80530178 0.000099976957 8053.017700023043 chA
000847 001271 001000 001840 036828 80540178 0.000099976918 8054.017700023082 chA
000848 001271 001000 001838 036803 80550178 0.000099976957 8055.017700023043 chA
000850 001271 001000 001838 036797 80560178 0.000099977062 8056.017700022938 chA
000849 001270 001000 001840 036838 80570178 0.000099977088 8057.017700022912 chA
000847 001271 001000 001838 036802 80580178 0.000099976902 8058.017700023098 chA
000848 001271 001000 001841 036848 80590178 0.000099976985 8059.017700023015 chA
000850 001271 001000 001840 036836 80600178 0.000099977086 8060.017700022914 chA
000849 001270 001000 001840 036831 80610178 0.000099977083 8061.017700022917 chA
000847 001270 001000 001839 036824 80620178 0.000099976970 8062.017700023030 chA
000849 001271 001000 001839 036819 80630178 0.000099977021 8063.017700022979 chA
000850 001272 001000 001840 036830 80640178 0.000099977028 8064.017700022972 chA
000849 001271 001000 001838 036813 80650178 0.000099977018 8065.017700022982 chA
000847 001270 001000 001839 036816 80660178 0.000099976965 8066.017700023035 chA
000848 001271 001000 001839 036820 80670178 0.000099976968 8067.017700023032 chA
000848 001270 001000 001840 036822 80680178 0.000099977023 8068.017700022977 chA
000849 001271 001000 001838 036817 80690178 0.000099977021 8069.017700022979 chA
000849 001270 001000 001839 036818 80700178 0.000099977075 8070.017700022925 chA
000848 001272 001000 001839 036818 80710178 0.000099976912 8071.017700023088 chA
000848 001271 001000 001839 036826 80720178 0.000099976972 8072.017700023028 chA
000849 001271 001000 001839 036827 80730178 0.000099977027 8073.017700022973 chA
000849 001270 001000 001840 036825 80740178 0.000099977079 8074.017700022921 chA
000848 001271 001000 001839 036823 80750178 0.000099976970 8075.017700023030 chA
000849 001271 001000 001840 036839 80760178 0.000099977034 8076.017700022966 chA
000847 001270 001000 001839 036819 80770178 0.000099976967 8077.017700023033 chA
000849 001271 001000 001841 036847 80780178 0.000099977038 8078.017700022962 chA
000849 001271 001000 001839 036818 80790178 0.000099977021 8079.017700022979 chA
000848 001270 001000 001838 036804 80800178 0.000099977012 8080.017700022988 chA
000849 001271 001000 001838 036806 80810178 0.000099977013 8081.017700022987 chA
000847 001271 001000 001840 036825 80820178 0.000099976916 8082.017700023084 chA
000849 001270 001000 001839 036818 80830178 0.000099977075 8083.017700022925 chA
000849 001270 001000 001841 036842 80840178 0.000099977090 8084.017700022910 chA
000847 001271 001000 001840 036834 80850178 0.000099976922 8085.017700023078 chA
000847 001271 001000 001838 036810 80860178 0.000099976907 8086.017700023093 chA
000849 001271 001000 001840 036834 80870178 0.000099977031 8087.017700022969 chA
000849 001271 001000 001840 036835 80880178 0.000099977031 8088.017700022969 chA
000848 001271 001000 001839 036826 80890178 0.000099976972 8089.017700023028 chA
000848 001271 001000 001838 036818 80900178 0.000099976967 8090.017700023033 chA
000849 001271 001000 001839 036822 80910178 0.000099977023 8091.017700022977 chA
000849 001271 001000 001839 036826 80920178 0.000099977026 8092.017700022974 chA
000849 001271 001000 001838 036803 80930178 0.000099977012 8093.017700022988 chA
000848 001271 001000 001839 036827 80940178 0.000099976972 8094.017700023028 chA
000847 001271 001000 001839 036819 80950178 0.000099976912 8095.017700023088 chA
000849 001270 001000 001839 036817 80960178 0.000099977075 8096.017700022925 chA
000850 001271 001000 001839 036826 80970178 0.000099977080 8097.017700022920 chA
000848 001271 001000 001839 036828 80980178 0.000099976973 8098.017700023027 chA
000848 001271 001000 001840 036847 80990178 0.000099976985 8099.017700023015 chA
000849 001271 001000 001839 036813 81000178 0.000099977017 8100.017700022983 chA
000848 001271 001000 001840 036840 81010178 0.000099976980 8101.017700023020 chA
000850 001271 001000 001840 036828 81020178 0.000099977081 8102.017700022919 chA
000849 001271 001000 001839 036826 81030178 0.000099977026 8103.017700022974 chA
000847 001270 001000 001840 036838 81040178 0.000099976979 8104.017700023021 chA
000849 001271 001000 001840 036834 81050178 0.000099977031 8105.017700022969 chA
000849 001271 001000 001840 036830 81060178 0.000099977028 8106.017700022972 chA
000849 001271 001000 001839 036827 81070178 0.000099977027 8107.017700022973 chA
000847 001271 001000 001839 036815 81080178 0.000099976910 8108.017700023090 chA
000848 001271 001000 001840 036822 81090178 0.000099976968 8109.017700023032 chA
000849 001271 001000 001838 036814 81100178 0.000099977019 8110.017700022981 chA
000849 001270 001000 001839 036821 81110178 0.000099977077 8111.017700022923 chA
000849 001270 001000 001839 036804 81120178 0.000099977066 8112.017700022934 chA
000848 001271 001000 001838 036805 81130178 0.000099976958 8113.017700023042 chA
000848 001271 001000 001840 036828 81140178 0.000099976972 8114.017700023028 chA
000849 001271 001000 001840 036833 81150178 0.000099977030 8115.017700022970 chA
000850 001271 001000 001839 036813 81160178 0.000099977072 8116.017700022928 chA
000849 001271 001000 001839 036822 81170178 0.000099977023 8117.017700022977 chA
000848 001272 001000 001840 036830 81180178 0.000099976919 8118.017700023081 chA
000849 001271 001000 001839 036829 81190178 0.000099977028 8119.017700022972 chA
000850 001271 001000 001839 036825 81200178 0.000099977080 8120.017700022920 chA
000849 001270 001000 001839 036810 81210178 0.000099977070 8121.017700022930 chA
000848 001271 001000 001840 036834 81220178 0.000099976976 8122.017700023024 chA
000848 001271 001000 001838 036818 81230178 0.000099976967 8123.017700023033 chA
000850 001271 001000 001840 036836 81240178 0.000099977086 8124.017700022914 chA
000849 001271 001000 001839 036827 81250178 0.000099977027 8125.017700022973 chA
000849 001271 001000 001839 036803 81260178 0.000099977011 8126.017700022989 chA
000847 001270 001000 001841 036841 81270178 0.000099976980 8127.017700023020 chA
000848 001271 001000 001839 036816 81280178 0.000099976965 8128.017700023035 chA
000849 001271 001000 001839 036816 81290178 0.000099977019 8129.017700022981 chA
000849 001271 001000 001840 036829 81300178 0.000099977027 8130.017700022973 chA
000849 001270 001000 001840 036829 81310178 0.000099977082 8131.017700022918 chA
000847 001271 001000 001839 036814 81320178 0.000099976909 8132.017700023091 chA
000848 001271 001000 001839 036818 81330178 0.000099976966 8133.017700023034 chA
000849 001271 001000 001838 036814 81340178 0.000099977019 8134.017700022981 chA
000850 001271 001000 001840 036829 81350178 0.000099977082 8135.017700022918 chA
000848 001271 001000 001837 036792 81360178 0.000099976950 8136.017700023050 chA
000848 001271 001000 001840 036832 81370178 0.000099976975 8137.017700023025 chA
000849 001270 001000 001839 036818 81380178 0.000099977075 8138.017700022925 chA
000849 001270 001000 001839 036815 81390178 0.000099977073 8139.017700022927 chA
000849 001270 001000 001840 036828 81400178 0.000099977081 8140.017700022919 chA
000847 001271 001000 001839 036826 81410178 0.000099976917 8141.017700023083 chA
000849 001270 001000 001840 036825 81420178 0.000099977079 8142.017700022921 chA
000849 001270 001000 001840 036833 81430178 0.000099977084 8143.017700022916 chA
000850 001271 001000 001840 036847 81440178 0.000099977094 8144.017700022906 chA
000847 001271 001000 001839 036826 81450178 0.000099976917 8145.017700023083 chA
000847 001271 001000 001840 036840 81460178 0.000099976926 8146.017700023074 chA
000849 001271 001000 001839 036817 81470178 0.000099977020 8147.017700022980 chA
000850 001271 001000 001839 036827 81480178 0.000099977081 8148.017700022919 chA
000849 001271 001000 001838 036812 81490178 0.000099977017 8149.017700022983 chA
000848 001271 001000 001839 036827 81500178 0.000099976972 8150.017700023028 chA
000849 001271 001000 001839 036823 81510178 0.000099977024 8151.017700022976 chA
000849 001271 001000 001840 036829 81520178 0.000099977027 8152.017700022973 chA
000849 001271 001000 001840 036825 81530178 0.000099977025 8153.017700022975 chA
000848 001271 001000 001838 036801 81540178 0.000099976956 8154.017700023044 chA
000848 001270 001000 001840 036833 81550178 0.000099977030 8155.017700022970 chA
000849 001271 001000 001839 036826 81560178 0.000099977026 8156.017700022974 chA
000850 001271 001000 001839 036818 81570178 0.000099977075 8157.017700022925 chA
000850 001270 001000 001839 036812 81580178 0.000099977126 8158.017700022874 chA
000848 001271 001000 001838 036798 81590178 0.000099976954 8159.017700023046 chA
000847 001271 001000 001838 036806 81600178 0.000099976905 8160.017700023095 chA
000849 001272 001000 001838 036802 81610178 0.000099976956 8161.017700023044 chA
000849 001271 001000 001838 036810 81620178 0.000099977016 8162.017700022984 chA
000848 001271 001000 001837 036786 81630178 0.000099976946 8163.017700023054 chA
000847 001271 001000 001838 036799 81640178 0.000099976900 8164.017700023100 chA
000847 001272 001000 001838 036817 81650178 0.000099976857 8165.017700023143 chA
000849 001271 001000 001840 036833 81660178 0.000099977030 8166.017700022970 chA
000849 001270 001000 001840 036826 81670178 0.000099977080 8167.017700022920 chA
000848 001271 001000 001837 036790 81680178 0.000099976949 8168.017700023051 chA
000848 001271 001000 001840 036837 81690178 0.000099976978 8169.017700023022 chA
000848 001271 001000 001840 036839 81700178 0.000099976979 8170.017700023021 chA
000849 001271 001000 001840 036835 81710178 0.000099977031 8171.017700022969 chA
000848 001271 001000 001839 036814 81720178 0.000099976964 8172.017700023036 chA
000847 001270 001000 001840 036828 81730178 0.000099976972 8173.017700023028 chA
000847 001271 001000 001840 036825 81740178 0.000099976916 8174.017700023084 chA
000848 001271 001000 001838 036802 81750178 0.000099976956 8175.017700023044 chA
000849 001271 001000 001840 036822 81760178 0.000099977023 8176.017700022977 chA
000848 001271 001000 001840 036834 81770178 0.000099976976 8177.017700023024 chA
000847 001270 001000 001840 036843 81780178 0.000099976982 8178.017700023018 chA
000848 001272 001000 001838 036814 81790178 0.000099976910 8179.017700023090 chA
000849 001271 001000 001839 036827 81800178 0.000099977027 8180.017700022973 chA
000850 001272 001000 001837 036793 81810178 0.000099977006 8181.017700022994 chA
000849 001271 001000 001839 036824 81820178 0.000099977025 8182.017700022975 chA
000848 001271 001000 001840 036827 81830178 0.000099976972 8183.017700023028 chA
000847 001271 001000 001839 036825 81840178 0.000099976916 8184.017700023084 chA
000848 001271 001000 001839 036819 81850178 0.000099976967 8185.017700023033 chA
000847 001271 001000 001839 036826 81860178 0.000099976917 8186.017700023083 chA
000847 001271 001000 001839 036810 81870178 0.000099976907 8187.017700023093 chA
000847 001271 001000 001839 036823 81880178 0.000099976915 8188.017700023085 chA
000848 001271 001000 001839 036813 81890178 0.000099976963 8189.017700023037 chA
000848 001270 001000 001838 036805 81900178 0.000099977013 8190.017700022987 chA
000848 001271 001000 001840 036830 81910178 0.000099976974 8191.017700023026 chA
000846 001271 001000 001839 036818 81920178 0.000099976857 8192.017700023143 chA
000847 001271 001000 001840 036831 81930178 0.000099976920 8193.017700023080 chA
000847 001271 001000 001839 036825 81940178 0.000099976916 8194.017700023084 chA
000848 001270 001000 001840 036824 81950178 0.000099977024 8195.017700022976 chA
000847 001271 001000 001839 036824 81960178 0.000099976916 8196.017700023084 chA
000845 001270 001000 001839 036810 81970178 0.000099976852 8197.017700023148 chA
000846 001271 001000 001839 036818 81980178 0.000099976857 8198.017700023143 chA
000847 001271 001000 001837 036777 81990178 0.000099976886 8199.017700023114 chA
000847 001271 001000 001839 036812 82000178 0.000099976908 8200.017700023092 chA
000847 001271 001000 001840 036830 82010178 0.000099976919 8201.017700023081 chA
000847 001271 001000 001838 036815 82020178 0.000099976910 8202.017700023090 chA
000847 001271 001000 001840 036846 82030178 0.000099976930 8203.017700023070 chA
000849 001270 001000 001840 036827 82040178 0.000099977080 8204.017700022920 chA
000847 001271 001000 001840 036828 82050178 0.000099976918 8205.017700023082 chA
000846 001271 001000 001839 036815 82060178 0.000099976855 8206.017700023145 chA
000847 001271 001000 001839 036824 82070178 0.000099976916 8207.017700023084 chA
000847 001271 001000 001840 036827 82080178 0.000099976917 8208.017700023083 chA
000847 001271 001000 001839 036820 82090178 0.000099976913 8209.017700023087 chA
000847 001270 001000 001839 036829 82100178 0.000099976974 8210.017700023026 chA
000847 001271 001000 001839 036829 82110178 0.000099976919 8211.017700023081 chA
000847 001271 001000 001840 036830 82120178 0.000099976919 8212.017700023081 chA
000847 001271 001000 001840 036823 82130178 0.000099976914 8213.017700023086 chA
000847 001271 001000 001839 036827 82140178 0.000099976918 8214.017700023082 chA
000847 001271 001000 001841 036843 82150178 0.000099976927 8215.017700023073 chA
000847 001270 001000 001839 036810 82160178 0.000099976961 8216.017700023039 chA
000849 001271 001000 001840 036829 82170178 0.000099977027 8217.017700022973 chA
000849 001271 001000 001838 036802 82180178 0.000099977011 8218.017700022989 chA
000850 001270 001000 001839 036818 82190178 0.000099977130 8219.017700022870 chA
000849 001271 001000 001839 036812 82200178 0.000099977017 8220.017700022983 chA
000849 001271 001000 001838 036807 82210178 0.000099977014 8221.017700022986 chA
000850 001271 001000 001838 036808 82220178 0.000099977069 8222.017700022931 chA
000849 001271 001000 001840 036834 82230178 0.000099977031 8223.017700022969 chA
000848 001271 001000 001839 036818 82240178 0.000099976966 8224.017700023034 chA
000847 001271 001000 001839 036827 82250178 0.000099976918 8225.017700023082 chA
000847 001271 001000 001838 036815 82260178 0.000099976910 8226.017700023090 chA
000849 001271 001000 001839 036826 82270178 0.000099977026 8227.017700022974 chA
000847 001271 001000 001839 036821 82280178 0.000099976914 8228.017700023086 chA
000847 001271 001000 001838 036800 82290178 0.000099976901 8229.017700023099 chA
000848 001271 001000 001839 036832 82300178 0.000099976975 8230.017700023025 chA
000849 001271 001000 001840 036829 82310178 0.000099977027 8231.017700022973 chA
000850 001271 001000 001838 036800 82320178 0.000099977064 8232.017700022936 chA
000849 001271 001000 001839 036818 82330178 0.000099977021 8233.017700022979 chA
000848 001271 001000 001839 036822 82340178 0.000099976969 8234.017700023031 chA
000848 001271 001000 001839 036818 82350178 0.000099976966 8235.017700023034 chA
000849 001271 001000 001838 036807 82360178 0.000099977014 8236.017700022986 chA
000849 001271 001000 001838 036804 82370178 0.000099977012 8237.017700022988 chA
000848 001271 001000 001839 036822 82380178 0.000099976969 8238.017700023031 chA
000847 001271 001000 001840 036837 82390178 0.000099976924 8239.017700023076 chA
000848 001270 001000 001838 036802 82400178 0.000099977011 8240.017700022989 chA
000849 001271 001000 001840 036832 82410178 0.000099977029 8241.017700022971 chA
000848 001271 001000 001839 036816 82420178 0.000099976965 8242.017700023035 chA
000847 001270 001000 001840 036817 82430178 0.000099976965 8243.017700023035 chA
000846 001271 001000 001840 036841 82440178 0.000099976872 8244.017700023128 chA
000848 001271 001000 001839 036822 82450178 0.000099976969 8245.017700023031 chA
000849 001272 001000 001837 036797 82460178 0.000099976954 8246.017700023046 chA
000847 001270 001000 001839 036813 82470178 0.000099976963 8247.017700023037 chA
000848 001271 001000 001841 036854 82480178 0.000099976989 8248.017700023011 chA
000848 001271 001000 001840 036840 82490178 0.000099976980 8249.017700023020 chA
000849 001271 001000 001839 036822 82500178 0.000099977023 8250.017700022977 chA
000848 001271 001000 001839 036819 82510178 0.000099976967 8251.017700023033 chA
000847 001271 001000 001839 036812 82520178 0.000099976908 8252.017700023092 chA
000848 001271 001000 001841 036845 82530178 0.000099976983 8253.017700023017 chA
000848 001271 001000 001838 036804 82540178 0.000099976958 8254.017700023042 chA
000849 001271 001000 001838 036816 82550178 0.000099977020 8255.017700022980 chA
000849 001271 001000 001840 036834 82560178 0.000099977031 8256.017700022969 chA
000847 001271 001000 001839 036825 82570178 0.000099976916 8257.017700023084 chA
000848 001271 001000 001840 036837 82580178 0.000099976978 8258.017700023022 chA
000849 001271 001000 001838 036805 82590178 0.000099977013 8259.017700022987 chA
000849 001271 001000 001841 036848 82600178 0.000099977039 8260.017700022961 chA
000848 001271 001000 001840 036827 82610178 0.000099976972 8261.017700023028 chA
000847 001271 001000 001837 036767 82620178 0.000099976879 8262.017700023121 chA
000848 001272 001000 001838 036810 82630178 0.000099976907 8263.017700023093 chA
000850 001271 001000 001839 036827 82640178 0.000099977081 8264.017700022919 chA
000849 001271 001000 001839 036827 82650178 0.000099977027 8265.017700022973 chA
000847 001271 001000 001839 036828 82660178 0.000099976918 8266.017700023082 chA
000846 001271 001000 001839 036826 82670178 0.000099976863 8267.017700023137 chA
000848 001271 001000 001838 036808 82680178 0.000099976960 8268.017700023040 chA
000848 001271 001000 001840 036834 82690178 0.000099976976 8269.017700023024 chA
000848 001271 001000 001840 036827 82700178 0.000099976972 8270.017700023028 chA
000848 001271 001000 001840 036834 82710178 0.000099976976 8271.017700023024 chA
000848 001271 001000 001840 036831 82720178 0.000099976974 8272.017700023026 chA
000849 001271 001000 001839 036816 82730178 0.000099977019 8273.017700022981 chA
000849 001271 001000 001840 036820 82740178 0.000099977021 8274.017700022979 chA
000848 001272 001000 001838 036811 82750178 0.000099976908 8275.017700023092 chA
000848 001271 001000 001839 036819 82760178 0.000099976967 8276.017700023033 chA
000848 001270 001000 001839 036816 82770178 0.000099977019 8277.017700022981 chA
000849 001271 001000 001838 036810 82780178 0.000099977016 8278.017700022984 chA
000849 001271 001000 001839 036829 82790178 0.000099977028 8279.017700022972 chA
000849 001271 001000 001840 036837 82800178 0.000099977033 8280.017700022967 chA
000848 001271 001000 001840 036827 82810178 0.000099976972 8281.017700023028 chA
000849 001271 001000 001839 036817 82820178 0.000099977020 8282.017700022980 chA
000850 001271 001000 001840 036826 82830178 0.000099977080 8283.017700022920 chA
000849 001270 001000 001839 036821 82840178 0.000099977077 8284.017700022923 chA
000847 001270 001000 001839 036816 82850178 0.000099976965 8285.017700023035 chA
000848 001271 001000 001839 036815 82860178 0.000099976964 8286.017700023036 chA
000849 001271 001000 001839 036823 82870178 0.000099977024 8287.017700022976 chA
000850 001271 001000 001840 036833 82880178 0.000099977084 8288.017700022916 chA
000849 001271 001000 001840 036826 82890178 0.000099977025 8289.017700022975 chA
000847 001271 001000 001840 036830 82900178 0.000099976919 8290.017700023081 chA
000849 001271 001000 001838 036800 82910178 0.000099977010 8291.017700022990 chA
000849 001271 001000 001840 036825 82920178 0.000099977025 8292.017700022975 chA
000850 001271 001000 001839 036824 82930178 0.000099977079 8293.017700022921 chA
000848 001270 001000 001840 036830 82940178 0.000099977028 8294.017700022972 chA
000847 001271 001000 001839 036815 82950178 0.000099976910 8295.017700023090 chA
000849 001271 001000 001839 036817 82960178 0.000099977020 8296.017700022980 chA
000849 001271 001000 001839 036805 82970178 0.000099977012 8297.017700022988 chA
000849 001271 001000 001839 036817 82980178 0.000099977020 8298.017700022980 chA
000848 001270 001000 001839 036816 82990178 0.000099977019 8299.017700022981 chA
000848 001271 001000 001840 036833 83000178 0.000099976975 8300.017700023025 chA
000849 001270 001000 001838 036812 83010178 0.000099977072 8301.017700022928 chA
000849 001271 001000 001840 036834 83020178 0.000099977031 8302.017700022969 chA
000848 001271 001000 001840 036839 83030178 0.000099976979 8303.017700023021 chA
000847 001270 001000 001839 036816 83040178 0.000099976965 8304.017700023035 chA
000849 001271 001000 001840 036838 83050178 0.000099977033 8305.017700022967 chA
000849 001271 001000 001840 036826 83060178 0.000099977025 8306.017700022975 chA
000849 001270 001000 001839 036825 83070178 0.000099977080 8307.017700022920 chA
000848 001271 001000 001839 036820 83080178 0.000099976968 8308.017700023032 chA
000848 001270 001000 001840 036829 83090178 0.000099977027 8309.017700022973 chA
000848 001271 001000 001839 036823 83100178 0.000099976970 8310.017700023030 chA
000849 001270 001000 001840 036832 83110178 0.000099977084 8311.017700022916 chA
000849 001270 001000 001839 036808 83120178 0.000099977069 8312.017700022931 chA
000847 001271 001000 001838 036815 83130178 0.000099976910 8313.017700023090 chA
000849 001271 001000 001839 036825 83140178 0.000099977025 8314.017700022975 chA
000848 001271 001000 001837 036794 83150178 0.000099976952 8315.017700023048 chA
000849 001272 001000 001838 036815 83160178 0.000099976965 8316.017700023035 chA
000849 001271 001000 001838 036800 83170178 0.000099977010 8317.017700022990 chA
000848 001271 001000 001839 036815 83180178 0.000099976964 8318.017700023036 chA
000849 001270 001000 001839 036813 83190178 0.000099977072 8319.017700022928 chA
000849 001271 001000 001839 036816 83200178 0.000099977019 8320.017700022981 chA
000849 001271 001000 001838 036809 83210178 0.000099977015 8321.017700022985 chA
000848 001271 001000 001839 036812 83220178 0.000099976962 8322.017700023038 chA
000848 001271 001000 001838 036800 83270178 0.000099976955 8327.017700023045 chA